
Multicahnnel temperature recording (test) instrument
An instruction manual
Vear1.0

Catalog

Chapter one     

Safety and warranty...................................

Introduction of instrument ...........................

Application  ..........................................

Chapter two     General overview of  temperature recorder ............. 

Difference of  temperature recorder ...................
Chapter three   fast introduction.....................................
Introduction of product front panel...................

Display and function key introduction.................
Product rear panel introduction.......................
Parameters setting....................................
Chapter four      Installation and accessories..........................

Chapter five       communication.........................................

                         Communication protocol................................ 

                         Software operation interface..........................

Chapter six        verification or calibration...........................

Preface

SHAPE  \* MERGEFORMAT



Summary

The new multi-channel temperature recorder has ultra-thin chassis, light weight, small volume and easy to carry.At the same time, the digital synchronous sampling technology has the advantages of fast measuring speed, high precision, easy to use, light and beautiful and so on. The instrument test precision is 0.5. This product is widely used in scientific research colleges and universities, electric power, chemical industry, petroleum, steelmaking, glass, ceramics, heat treatment and plastic, rubber, printing and dyeing, packaging and food machines. Temperature measurement and temperature control in the machinery industry.
The disposal of old packaging boxes and the disposal of old packaging boxes should be carried out in accordance with relevant national environmental regulations.

Please pay attention to separate the plastic and paper packaging products.

Maintenance of the instrument

SHAPE  \* MERGEFORMAT



◆ The surface of instrument should be dust in time,The interior of the instrument is best not to accumulate dust.

◆ The instrument should be calibrated once a year, so as to ensure the accuracy of the instrument, so that it can be used normally.

Warning: be sure to disconnect the power before cleaning.

Announcement

The contents of this manual are subject to change without notice. Excuse me!

Chapter one

Security

Do not install on your instrument instead of parts, or execute any unauthorized modification.Please send the instrument to the maintenance department of the company to maintain it, so as to ensure its safety characteristics.

Security rules

In order to prevent electric shock, not authorized personnel of the company, it is strictly forbidden to disassemble the instrument and strictly prohibit the use of the equipment on the life maintenance system or any other equipment with safety requirements.We are not responsible for direct or indirect financial losses that may occur when using this product.

Certification and quality assurance

The multichannel temperature recorder has achieved all the technical indicators mentioned in the manual.

warranty

The company of materials and manufacturing of this product, from the date of shipment to the quality warranty for one year.

Warranty service

If you need the product warranty service or repair, must be the product to the maintenance unit. The customer must pay the one-way freight to the maintenance department, and the company will be responsible for paying the return freight. If the product is returned to factory maintenance from other countries, all freight, duties and other taxes shall be borne by the customer.

Guarantee restriction

The warranty does not apply to the damage caused by the following circumstances:

The customer is incorrect or improper in repairing products.

Customers use their own software or interface;

Unauthorized modification or misuse;                                                       

Operating this product or not in the specified environment or repair and repair at unsuitable places,Damage caused by the circuit of the customer's own installation.Or a defect caused by a customer's use of its own product.

Brief introduction 

JK3000 is a high performance, low price temperature recorder. It can watch multi-channel temperature changes at the same time.It is very suitable for application of temperature acquisition and temperature recording.It has the RS232 interface to upload data directly to the PC, and also has the USB interface to save the measured data directly to the U disk when it is inconvenient to connect to the PC, and then transfer the data to the PC when needed.It is a half frame wide host with a top 64 channel direct slot on the back of which can be measured with a thermocouple, and 1 infinite expansion slots, which can be assembled by a module.Whether you need only a few simple data recording channels or hundreds of thousands of performance channels, you can meet your temperature collection at a reasonable price.
3 different types of thermocouple measurement

Measurements that can be dependent on the temperature recording software

The unique design allows every 8 channels to be configured for a single unit, so as to achieve maximum flexibility and quick and convenient setting of automatic range conversion inside. That is to say, choose 1-8 as K type thermocouple, 9-16 as J type thermocouple, and change configuration according to needs.

Here are 64 direct access slots and 1 infinite expansion slots on the back to meet your unique needs. You can only buy the parts you need at that time, and then add more modules as your application develops. 1.5 rack units can be used to measure up to hundreds of thousands of inputs.

Free software simplifies data acquisition

If we require data recording capability based on PC, but we don't want to spend time programming, software can solve this problem. With this software, you can set up your test, collect and archive, manage and measure data, and display and analyze the data and curves in real time. The familiar electronic table environment can be easily configured and controlled for testing. Rich color graphics provide many options for displaying your data. In addition, using software can easily send data to other applications for further analysis, or include it in your display and report.

A brief introduction to the module

The module can be directly used to connect the infinite expansion slot to the back of the instrument and connect the thermocouple line.

When you need more modules, you only need to connect to the reserved slot of the first module.(this function is especially practical when the instrument and the need to be far away from the measured point, not only can reduce the measurement accuracy caused by too long thermocouples, but also save the large amount of money cost needed for the purchase of thermocouples).

application

Instrument safety index

SHAPE  \* MERGEFORMAT



◆Insulation resistance: the outer part of the housing and the power input are larger than 20MΩ.
◆ Withstand voltage: the power input terminal and the housing are 2 seconds withstanding voltage 1800V, and the rated current is 10mA.
◆Grounding: the power grounding pole, and the resistance between the outer shells is less than  0.2Ω。

◆ security: conforms to  CSA,UL-1244,IEC1010 Cat I  regulations  RFI  and ESD:CISPR11,IEC801/2/3/4

Precautions before use of the instrument

SHAPE  \* MERGEFORMAT



◆ The instrument should be placed smoothly And where the dust is less.

◆ The temperature conditions that ensure the accuracy of the test:20℃±5℃。

◆ Working temperature range:0~40℃，Air relative humidity：30%~90%。

◆ No corrosion, detonable gas in the room.

◆ Protect ground.

◆ Power supply: AC90-240V/50/60HZ，Power waste ≤15W。

◆ Fuse：is  0.5A

◆ weight ：Net weight  3.6kg

◆ accuracy :±0.3%(the highest temperature measured at the time)

When using the "T" thermocouple, the accuracy can reach as high as 0.1%.

Chapter two

A list of functional difference of each thermometer

	Serial number
	Model 
	The largest number of channel
	temperature range①
	U DISK 
②
	RS232 communication interface

	1
	JK3000-8
	8
	3 kinds of thermocouples
	yes
	yes

	2
	JK3000-16
	16
	3 kinds of thermocouples
	yes
	yes

	3
	JK3000-24
	24
	3 kinds of thermocouples
	yes
	yes

	4
	JK3000-32
	32
	3 kinds of thermocouples
	yes
	yes

	5
	JK3000-40
	40
	3 kinds of thermocouples
	yes
	yes

	6
	JK3000-48
	48
	3 kinds of thermocouples
	yes
	yes

	7
	JK3000-64
	64
	3 kinds of thermocouples
	yes
	yes


The specification and measurement range of the S series input type

	Input type
	Specifications
	measuring range

	Thermocouple
	J
	-200~1200℃

	
	K
	-200~1370℃

	
	T
	-200~400℃


technical parameter ：

	Model                             
	multichannel temperature tester

	Input type                      
	Thermocouple：J/K/T

	measuring range        
	Measuring range：-200℃～1760℃      J,K,T,
measurement accuracy： 0～1000℃:±(reading value×0.5%+1)℃，-100～0℃:±(reading value×0.5%+2)℃；

	Number of channels                        
	8，16，24，32，40，48，64 channels

	Display                                   
	480*272  5 inch TFT color screen

	Display resolution                                                                
	1mV  0.1℃

	U disk storage                                                                    
	yes

	scanning speed                                                                    
	100ms

	Internal storage                                                                     
	yes

	Curve portrayal                                                                   
	yes

	communication interface                                                                
	RS232 or USB

	software kit                                                               
	2015 version of V1.3 software

	Alarm mode                                                           
	One way public alarm (relay output).(selection)

	Voltage output                                                              
	Internal independent power supply output

	Anti-interference                                                                    
	Anti high frequency interference

	Single screen display                                                                    
	32 channel

	Temperature correction                                                                    
	yes

	Other functions                                                             
	Clock function ,calendar function

	Terminal type
	M3 screw type

	Permissible environmental conditions
	Power supply：AC 220V±10%，50Hz±2%；   

2、Working environment: working temperature:-20－70℃，relative humidity：20%－90%；

	Measuring range
	J type:       -50.0℃~ 1000.0℃   

K type:       -50.0℃~ 1200.0℃  

T type:       -50.0℃ ~ 400.0℃  

	attachment
	Standard: K type thermocouple 8 (2 meters / root) data acquisition software.

	Size 
	Length * width * height: 36cm * 26cm * 16cm weight: About 5kg


3 kinds of thermocouples in the table ① ,3 different types of thermocouples The range of temperature is：

	J  thermocouple
	0 ～ 760℃

	K thermocouple
	0 ～ 1370℃

	T thermocouple
	-100 ～ 400℃


U disk: the instrument does not connect to PC. It can store data on the U disk, and save data permanently in the computer by transferring data.

The 3000-8 in Table 8 is that the acquisition module with 8 channels can be used in conjunction with each of the above instruments. After the customer has purchased any instrument, if you want to increase the acquisition channel, the 8 channel can choose 108 modules for one unit.

In table 10, 3000+108 indicates that customers can choose how many channels to collect temperature according to their needs and choose a 108 acquisition module with 8 channels as a unit in the case of a host. If you need 80 channels, you can choose a 64 way host and 2 108 acquisition modules.

Chapter three

Product introduction


front panel

[image: image5.jpg]front panel

QS 00Q00o

s





Introduction of front panel

1  ：display screen
2  ：power 
3 ：System function key
4  ：USB disk interface
5  ：Qualified / unqualified

6  ：Lock key
7  ：Digital keyboard

8：Cursor key

9：show key

SHAPE  \* MERGEFORMAT



Display area schematic:

[image: image7.jpg]<Measure *
001
002
003
004
005
Q06
Qo7
008

18-0%-19
10:35:05





1 ：Show the temperature of each channel, respectively（001 represents first channels, 002 represents second channels), and so on.

2 ：Display date / time: as shown in the picture  1th January,2007,At 20:26 00 seconds

3：Thermocouple model

4: Unit of temperature

Rear panel schematic:

[image: image8.png]



As shown in the picture.

1：Indicating the interface definition of each channel number and input signal（01 represents 1 channels, 02 represents 2 channels, and so on.）

Signal wiring method：

[image: image9.png]CB

A\CBA

CBA| cBAl





Thermocouple        Thermal resistance      voltage        current                         
Signal access       Signal access          Signal access    Signal access

2：The input voltage of the power outlet is  AC220V
 50Hz  0.5A

3：The USB interface is used for computer USB port connection (RS232 and USB choose one).

4：The RS232 interface can directly upload the data measured to the PC through the RS232 cable line and the host computer software (RS232 and USB choose one).

5:CASCADING: Used for channel expansion and alarm relay contact output terminal, when alarm relay output, contact signal as shown in the figure:

[image: image10.png]



2,3 are normally closed contacts

3,4 are normally open contacts

Contact load capacity is DC 24V/1A

Channel setting ：

Select the required setup items, press the "parameter settings" key, and use the "up and down" key to set the value or function.For example, setting the number of channels: after moving the cursor to the number of channels, press the key of "parameter", change the parameters with "up and down" key, then press the "set to finish" button.

Type setting of thermocouple: move cursor to the type of thermocouple, select thermocouple type with "up and down" key, and press "set up" to exit.

The other projects are like this,

When the settings are completed, press "MENU" to save the settings and exit; press the exit key, do not save the settings and exit.

[image: image11.jpg]/CHH | ,- G

BUZEER ﬁLﬁBNOFF,M .
High Alarm +0000(},; .
CJC Correct CHN 11- 205 4.; .
TIME SET  2018.05.19  10:35: 3__.,__,,,,__7?“7 ’
THERM TYPE CHN 11-20  TYPE:TC-J .

150019
10:35:34





Equipment parameter page introduction:

Channel number: number of test instruments

Baud rate: communication rate between instrument and host computer

Device address: the host computer inquires the address of the machine

Sampling time: speed of reading temperature and data

Unit of temperature: a unit of temperature:（℉，℃）

Automatic screen change: when the number of passages exceeds 32, the temperature needs multi page display.

Multi page automatic rotation display at ON

Temperature rise: no function at present

Time setting: instrument clock

Buzzer alarm: the buzzer opens or closes the choice

Power saving: no function at present

Upper limit alarm value: high temperature alarm value. Every 8 channels is an adjustment base, and can be switched through thermocouple type.                

The lower limit alarm value is low temperature alarm value. Every 8 channels are adjusted base and can be switched through thermocouple type.

Reference value: no function at present

CJC calibration: adjust the channel deviation value, each 8 channels as an adjustment base.

Type of thermoelectric coupler: Type: choose the probe type of the instrument test, each 8 channels as an adjustment base.

Functional selection:

1. When displaying the channel parameters, press the "MENU" button to switch to the digital bar.Then, through the upper and lower keys to adjust the number you need, and then switch to the next digit through the left and right keys.When the settings are completed, press "MENU" to save the settings and exit; press the exit key, do not save the settings and exit.。

2. Baud rate： 1200，2400，4800，9600，19200，38400 According to your different needs, you can set it on your own. The default is 9600 when you start the machine.

3. Device address: according to the COM port of different PC, that is to choose to several. The default is 001.

4. Sampling time: how long is it needed to collect, how long it takes to tune up, the minimum time is 1 second, the maximum is 9999 seconds.

5. Temperature unit：℃/℉ Switch.

6. Automatic screen switching：ON/OFF Switch.

7.Time settings: set the year, month, hour, minute and seconds that you need to display.

8. Buzzer alarm: when the temperature is adjusted to "ON", if the temperature exceeds the upper limit or lower limit set by you, there is a sound of sound. If it is adjusted to "OFF", there is no sound hint.

9. In the measurement, you can set an upper limit of temperature and a lower limit of a temperature. When the measured temperature exceeds the range you set, if the buzzer alarm is in the "ON" state, there will be a sound warning to you. The "+" on the display represents the positive temperature "-" representing the negative temperature.
10. Type of thermocouple: it can be arranged from 8 to one, and different thermocouple types are set for a unit.

11. After all settings are finished, press "set up" to exit the parameter settings.

Curve parameter page introduction:

At any interface, just press the "DISP" button to return to the measurement display interface.

According to the "graph" key, the curve display interface is displayed, as shown in the figure.

[image: image12.jpg]+1500 |
+1312 |
+1125 |
+0937 |
+0750 |
+0562 |
+0375 |
+0187 |
| +0000,





According to the "curve setting" key, enter the "curve setting" display interface, as shown in the figure.

[image: image13.jpg]<CURV SET>

X_Lendth ey 0
xeetd s
BGCLR  o0a fili

TO1Curve CLRO10 [FF
T03Curve CLROT1 [T |
T05Curve CLR130 [l ' TO6Curve CLR 160 e
TO7Curve CLR190 | T08Curve CLR 220 BN
CHN DISP  CHN 01-08 11111111  Font CLR 003 I

1805019
10:36:41





X axis length: X axis time value, 1min~999min can be set freely.

Y axis range: Y axis temperature range, -1999~+1999,If you want to set negative numbers, you need to make sure that when the number is not 0 Change+/-.
X axis grid number: X axis grid number, 1~9 can be set freely.

Y axis grid number: Y axis grid number, 1~9 can be set freely.

Background color: curve background color value, 000~255 can be set freely.

Raster color: curve grid line color, 000~255 can be set freely.

Font color: the color of the word on the curve page, 000~255 can be set freely.

Curve color: set the color value of each channel curve, 000~255 can be set freely.

Channel display: channel curve opening and closing settings, 0 passes, 1 open, every 8 channels is a set base.
Introduction of function settings interface

At any interface, just press the "DISP" button to return to the measurement display interface.

Press the function setting key to enter the function setting interface, as shown in the figure.

[image: image14.jpg]FUNC SET

Language hinese Enyliae q,
Sav mode U save

18-05-19
10:37:20





Language: language includes two languages, Chinese and English.

Set steps: the cursor area moves to Chinese and English, click the "parameter setting" key, switch the language, click "set up" key, if you need to cancel the setting, click "cancel the setting" key.

Storage mode: U disk storage and internal storage of 2.
Introduction of Function setting interface

At any interface, just press the "DISP" button to return to the Function setting  interface.

According to the U disk , you can enter the Function setting interface, as shown in the figure.

[image: image15.jpg]Lanfuafe  Chinese EnSlish
Sav wode

18-05-19
10:55:56




Press the up and down button to select the U save  and click Set End .
system information
[image: image16.jpg]DEVCE TPYE —0000
SOFT VER Yer:0.0
HARDUARE VER  Ver:0.0
PRODOUCT DATE 0000-00-00
DEVCE 1D 18A000000




The system information interface includes:

Instrument model

Software version

Hardware version

Date of manufacture

Instrument number

Chapter four

	Power line (one)
	JK-E171 JK-E172 JK-E173

JK-E174
	Users can choose different power lines according to the specifications of the power outlets in the area.
	one

	User's Manual
	
	Including the installation of information, operation information
	one

	Software CD-ROM
	
	Software and technical information
	one

	Test Report
	
	Test report on the instrument before the factory
	one

	U disk
	
	
	

	RS232 Cable
	JK-R1
	
	one


install

Make sure that you receive the following components when you receive multiple temperature records. If there is any shortage, please contact your dealer as soon as possible.

Chapter five: instrument communication
Communication interface reference for multichannel temperature recorder
To facilitate users, there are two optional interfaces for instrument communication: RS232 interface (DB9 port) and USB port.

Baud rate: range 1200 - 38400

Instrument communication data format: 1 start bits, 8 data bits, 1 stop bits (corresponding to 8051, 8096 serial mode 1).

Serial connection:Standard RS 232 level, DB9 pin definition,2 - RXD, 3 - TXD, 5 - GND

USB communication interface: the inside is a 232 turn USB module, and the outlet is located on the back side of the instrument (the front panel is the U disk socket).

USB communication requirements: we need to install USB driver software on the U disk to be used.

Connect with the computer: connect the USB communication line with the instrument to the computer USB port. After the USB driver is installed, the computer will detect the USB communication interface of the instrument. The resource manager can check the port number, and the port number of the communication software can be communicated to the secondary port number.
Instrument communication command:

Start a real time acquisition handshake ：01  11  c0  2c

01 Instrument address，11 command，C02c is the CRC check value.
Instrument return：01  11  0a  49  44  41  51  38  35  31  38  30  30  5b  d6

5bd6 is a CRC check value

Reading upper and lower limits：01  03  00  50  00  03  05  da

05da is a CRC check value

Instrument return： 01  03  06  00  01  00  32  00  0b  fc  bd

0001 represents the first set of parameters, the upper limit of 0032, and the lower limit of 00 0b.fcbd is a CRC check value

3,Read the data：01  03  00  00  00  50  45  f6

50 Data length，45 f6 is a CRC check value

return：01 03 a0   00b4 00b5  00b3………………CRC

The data is 2 bytes with a symbol number, a decimal point

4，Write set data 01  10  00  50  00  03  06  00  01  00  32  00  0b  38 09(CRC)

The 03 is 3 sets of data, 06 is 6 bytes, 0032 is the upper limit value, the 000b lower limit value.

The data is 2 bytes with symbols, no decimal point.

Software operation interface

1, after ensuring the connection of RS232 communication cables between PC and instruments, After installing the software provided by our company, click on the JK5080 icon and you will enter it. In the "File" column, click on the "[image: image17.png]


" Select the save path:

[image: image18.png]Language

]

u

@ u L]

° 06 |0 | m u
o7 u u

T | W | w | m

| o | m

0 u ]
W o | m

3 & 130 180 20 300 350 0 450 520 600 660 730 780 80 %0 %0 1020 180 140 100 015 : g

Measure Data

Disconnect 2019-09-03 08:33:05





2. Click on the communication settings option to refresh the serial password. The baud rate of 9600 remains unchanged.

[image: image19.png]-
o
.
M

720 1020

Measure Data

Disconnect




Click on the instrument settings option to select the thermocouple model.

[image: image20.png]INKOZSRBEREERNVI — 2HE

50 []
4 —001 [ 001 []
40

o €3 002 034 ]
30 003 1™ 003 ]
2 0% 06| W

20 — 005
07 | m

15 - 006
006, []

N 007
31 o 00 | W
020406080 110 140 170 200 230 260 230 350 350 380 410 440 470 500 530 560 530 620 630 680 710 740 770 800 830 830 850 930 950 950 1020 1060 1100 1140 1180 []

NEHE
== B BiE) ch1 ch2 ch3 Cha chs Che ch7 chs

2019-07-27 14:40:27




Click on the choice of curve settings, select the upper and lower limits.

[image: image21.png]JINKO JinAiLian temperature acquisition recorder(SerialPort) o x
Commsettings g | splay  Language
Thermocoup
= B
001 []
P 002 []
003 []
004 0 | m
— o001 07 | m
0 ooz [ []
o003 00 | m
004 [ o0g 0 | W
&= u o
0 =C3 009 u ]
o7 (] []
T | W | w | m
| o [ m
-0 (] []
m | o5 | m
0 60 120 180 240 300 360 420 480 540 600 660 720 780 240 900 960 1020 1080 1140 1200 o5 L _ u .
e ™ ™
Measure Data v
Number Date Time Chl Ch2 Ch3 Cha Chs Ché Ch7 Ch8

Disconnect 2019-09-03 08:34:13




Click on the measurement display, click on the start of the test, a curve appears.

[image: image22.png]Language

Disconnect

[]
001 []
P 002 []
003 []
0 | m
— o001 07 | m
0 ooz [ []
o003 00 | m
004 [ o0g 0 | W
&= u o
0 =C3 009 u ]
o7 (] []
T | W | w | m
| o [ m
-0 (] []
m | o5 | m
0 60 120 180 240 300 360 420 480 720 0L u _ u
e ™ ™
Measure Data
Number Date Time Chl Ch2 Ch3 Cha Chs Ché Ch7 Ch8

2019-09-03 08:35:29





In the measurement display interface, click on the [image: image23.png]=4

Testing



testing, select the Excel file in the save path and view the historical data.

[image: image24.png]=4

Testing





In the File Options Interface, click View to view historical data and graphs in the save path, as well as data stored on the U disk from the instrument.

[image: image25.png]Languag

Measure Result

720 1020

Measure Data

Disconnect 2019-09-03 08:36:55





Wireless mode transmission:

Supporting virtual serial port, the computer must have wireless receiving function, and install virtual serial port software. After installation, open virtual software and add a serial port.

[image: image26.png]R BAEHISORE V371520
BE0) T81) HEAO) Englsh #EhH)

. KoL A R COM2 -

2240 B0 9

P N = . v Rt [TOPClent -
| Wik | EE |Suitd HiF | #% BihalE

808 ERE= BORE Rt

EfPAE:  [19216811254

il et
FERO 234

&

Quwih | Qnw| =z





Set up the back point to determine.

The instrument is manufactured in wireless AP (TCP server) mode with wireless name:JK5000,
IP :192.168.11.254

If a laptop is connected to an instrument, the network connection of the laptop should be changed to the JK5000 network in order to achieve wireless communication.

Open the temperature acquisition software, change COM port to virtual serial port in the settings, and then data can be processed.












